## information
Electronic supplementary material for:
Rapid quantification method of pharmaceuticals in urban wastewater using UHPLC-Orbitrap mass spectrometry – a multi target approach for EU-directive advanced treatment evaluation
Journal: Analytical and Bioanalytical Chemistry
Uta Hofmanna, Reinhard Oertela, Nancy Engela, Johannes Zanderb, Annemarie Lipperta, Sara Schubertb, Bertold Rennera
corresponding author: Uta Hofmann, uta.hofmann@tu-dresden.de
TUD Dresden University of Technology, Institute of Clinical Pharmacology, Faculty of Medicine Carl Gustav Carus, Dresden, Germany
TUD Dresden University of Technology, Institute of Hydrobiology, Dresden, Germany

## regression and Q-limits
SI-1: all calibration data in three batches (KARL 1 - KARL 3) with six standard curves |  |  |  |  | SI-2: calibration data means over three batches |  |  |  | SI-3: LOD, LOQ and ULOQ
compound | y_intercept | slope | r_2 | sample_id | compound | mean_slope | mean_y_intercept | mean_r2 | compound | blank_max_height | blank_mean | slope | lod_concentration | LOQ_mean | ULOQ_mean | ULOQ_median | n
Amisulpride | -0.007212 | 0.001308 | 0.9991 | KARL_1 | Amisulpride | 0.001267 | -0.008033 | 0.9984 | Amisulpride | 43423.5342 | 0.0056184 | 0.001488 | 3.508 | 16.6667 | 15000 | 20000 | 6
Metoprolol | 0.0008349 | 0.0001079 | 0.9993 | KARL_1 | Candesartan | 0.00061697 | -0.010572 | 0.99763 | Candesartan | 2067.5559 | 0.0021549 | 0.0004428 | 16.6473 | 62.5 | 11250 | 10000 | 6
Tramadol | 0.002766 | 0.0003676 | 0.9984 | KARL_1 | Carbamazepin | 0.00079787 | 0.0029231 | 0.99823 | Carbamazepine | 48421.1333 | 0.0033801 | 0.0008436 |  | 15 | 13750 | 15000 | 6
Venlafaxine | 0.008065 | 0.0005316 | 0.9991 | KARL_1 | Citalopram | 0.0012853 | -0.0095947 | 0.99807 | Citalopram | 15936.4908 | 0.00049849 | 0.001541 | 2.3361 | 25 | 12500 | 10000 | 6
Propranolol | 0.01531 | 0.0008493 | 0.999 | KARL_1 | Clarithromycin | 0.001895 | -0.008482 | 0.9979 | Clarithromycin | 19563.5957 | 0.00052609 | 0.001468 | 2.1599 | 25.8333 | 8333.3333 | 10000 | 6
Citalopram | -0.005728 | 0.001386 | 0.9971 | KARL_1 | Diclofenac | 0.00058828 | -0.12404 | 0.99633 | Diclofenac | 1353.5665 | 5.1897e-05 | 3.055e-06 | 100.3906 | 325 | 23500 | 25000 | 5
Carbamazepin | 0.005132 | 0.0008105 | 0.9982 | KARL_1 | Irbesartan | 0.0014633 | -0.028645 | 0.99847 | Irbesartan | 0 | 0 | 0.001461 | 7.3924 | 53.3333 | 14583.3333 | 20000 | 6
Clarithromycin | -0.002145 | 0.001935 | 0.9967 | KARL_1 | Metoprolol | 0.00010777 | 0.00019859 | 0.99917 | Metoprolol | 24166.7031 | 0.00056436 | 0.0001356 |  | 28.3333 | 17083.3333 | 20000 | 6
Candesartan | -0.005016 | 0.0006449 | 0.9984 | KARL_1 | Propranolol | 0.00080763 | 0.0037652 | 0.99877 | Propranolol | 78538.5991 | 0.0083001 | 0.001001 | 3.8009 | 17.5 | 9166.6667 | 7500 | 6
Irbesartan | -0.008056 | 0.001554 | 0.9992 | KARL_1 | Tramadol | 0.00035153 | 0.00019873 | 0.99827 | Tramadol | 22689.4273 | 0.0019571 | 0.0003959 | 10.8626 | 28.3333 | 14166.6667 | 15000 | 6
Valsartan | -0.003916 | 0.0009234 | 0.9982 | KARL_1 | Valsartan | 0.0008535 | -0.007815 | 0.99773 | Valsartan | 2786.2846 | 0.01244 | 0.000126 | 151.098 | 30 | 7083.3333 | 15000 | 6
Diclofenac | 0.0005507 | 1.478e-05 | 0.9965 | KARL_1 | Venlafaxine | 0.00050897 | 0.0033052 | 0.99907 | Venlafaxine | 45444.0613 | 0.0019348 | 0.000428 |  | 20 | 16666.6667 | 20000 | 6
Amisulpride | -0.008248 | 0.00121 | 0.9978 | KARL_2
Metoprolol | -0.0003116 | 0.0001075 | 0.9991 | KARL_2
Tramadol | -0.002484 | 0.0003427 | 0.998 | KARL_2
Venlafaxine | 0.0017 | 0.0004821 | 0.9994 | KARL_2
Propranolol | -9.146e-05 | 0.0007553 | 0.9994 | KARL_2
Citalopram | -0.01796 | 0.001243 | 0.9987 | KARL_2
Carbamazepin | 0.003769 | 0.0007423 | 0.9976 | KARL_2
Clarithromycin | -0.01572 | 0.00192 | 0.9994 | KARL_2
Candesartan | -0.01562 | 0.0006002 | 0.9984 | KARL_2
Irbesartan | -0.04188 | 0.001387 | 0.9986 | KARL_2
Valsartan | -0.01014 | 0.0008251 | 0.996 | KARL_2
Diclofenac | -0.00167 | 1.205e-05 | 0.9952 | KARL_2
Amisulpride | -0.008639 | 0.001283 | 0.9983 | KARL_3
Metoprolol | 7.248e-05 | 0.0001079 | 0.9991 | KARL_3
Tramadol | 0.0003142 | 0.0003443 | 0.9984 | KARL_3
Venlafaxine | 0.0001506 | 0.0005132 | 0.9987 | KARL_3
Propranolol | -0.003923 | 0.0008183 | 0.9979 | KARL_3
Citalopram | -0.005096 | 0.001227 | 0.9984 | KARL_3
Carbamazepin | -0.0001316 | 0.0008408 | 0.9989 | KARL_3
Clarithromycin | -0.007581 | 0.00183 | 0.9976 | KARL_3
Candesartan | -0.01108 | 0.0006058 | 0.9961 | KARL_3
Irbesartan | -0.036 | 0.001449 | 0.9976 | KARL_3
Valsartan | -0.009389 | 0.000812 | 0.999 | KARL_3
Diclofenac | -0.371 | 0.001738 | 0.9973 | KARL_3

## specificity accuracy precision
SI-4: interday precision and accuracy | [5 empty] | SI-5: intraday precision and accuracy
compound | theoretical_amount | Mean_all | SD_all | CV_inter | AC_inter | compound | sample_id | theoretical_amount | Mean_intra | SD_intra | CV_intra | AC_intra
Amisulpride | 400 | 416.9083 | 29.3918 | 7.0499 | 104.2271 | Amisulpride | KARL_1 | 400 | 388.8763 | 28.3217 | 7.2829 | 97.2191
Amisulpride | 2000 | 1914.9312 | 135.1858 | 7.0596 | 95.7466 | Amisulpride | KARL_1 | 2000 | 1788.1873 | 73.2076 | 4.094 | 89.4094
Amisulpride | 4000 | 4087.8504 | 183.8194 | 4.4967 | 102.1963 | Amisulpride | KARL_1 | 4000 | 3878.3407 | 77.0437 | 1.9865 | 96.9585
Candesartan | 400 | 403.6133 | 44.6515 | 11.0629 | 100.9033 | Candesartan | KARL_1 | 2000 | 1660.791 | 72.4654 | 4.3633 | 83.0395
Candesartan | 2000 | 1806.2639 | 149.5066 | 8.2771 | 90.3132 | Candesartan | KARL_1 | 4000 | 3508.3845 | 90.6512 | 2.5838 | 87.7096
Candesartan | 4000 | 3731.0727 | 256.3671 | 6.8711 | 93.2768 | Candesartan | KARL_1 | 400 | 349.1951 | 8.8733 | 2.5411 | 87.2988
Carbamazepin | 400 | 429.1251 | 19.2127 | 4.4772 | 107.2813 | Carbamazepin | KARL_1 | 400 | 431.5214 | 28.8451 | 6.6845 | 107.8803
Carbamazepin | 2000 | 1863.2001 | 75.3541 | 4.0443 | 93.16 | Carbamazepin | KARL_1 | 2000 | 1841.6573 | 58.6427 | 3.1842 | 92.0829
Carbamazepin | 4000 | 3828.247 | 190.9444 | 4.9878 | 95.7062 | Carbamazepin | KARL_1 | 4000 | 3784.8464 | 74.5676 | 1.9702 | 94.6212
Citalopram | 400 | 417.5783 | 43.1914 | 10.3433 | 104.3946 | Citalopram | KARL_1 | 4000 | 3641.2959 | 164.0626 | 4.5056 | 91.0324
Citalopram | 2000 | 1865.4627 | 140.0241 | 7.5061 | 93.2731 | Citalopram | KARL_1 | 400 | 368.7532 | 14.3513 | 3.8918 | 92.1883
Citalopram | 4000 | 4098.5971 | 406.7677 | 9.9246 | 102.4649 | Citalopram | KARL_1 | 2000 | 1724.056 | 55.5102 | 3.2197 | 86.2028
Clarithromycin | 400 | 453.849 | 29.7834 | 6.5624 | 113.4623 | Clarithromycin | KARL_1 | 400 | 448.0075 | 21.9703 | 4.904 | 112.0019
Clarithromycin | 2000 | 1861.6565 | 84.4001 | 4.5336 | 93.0828 | Clarithromycin | KARL_1 | 2000 | 1861.5085 | 81.8962 | 4.3995 | 93.0754
Clarithromycin | 4000 | 4122.1878 | 232.9631 | 5.6514 | 103.0547 | Clarithromycin | KARL_1 | 4000 | 4062.1683 | 80.5582 | 1.9831 | 101.5542
Diclofenac | 2000 | 2071.368 | 68.4647 | 3.3053 | 103.5684 | Diclofenac | KARL_1 | 20000 | 18602.1652 | 1392.3501 | 7.4849 | 93.0108
Diclofenac | 10000 | 9988.7068 | 445.7379 | 4.4624 | 99.8871 | Diclofenac | KARL_1 | 10000 | 10241.1533 | 312.611 | 3.0525 | 102.4115
Diclofenac | 20000 | 19822.331 | 1270.1946 | 6.4079 | 99.1117 | Diclofenac | KARL_1 | 2000 | 1999.3983 | 39.0773 | 1.9545 | 99.9699
Irbesartan | 400 | 429.4189 | 32.5857 | 7.5883 | 107.3547 | Irbesartan | KARL_1 | 4000 | 3677.4452 | 363.8536 | 9.8942 | 91.9361
Irbesartan | 2000 | 1955.3095 | 213.4417 | 10.916 | 97.7655 | Irbesartan | KARL_1 | 2000 | 1761.75 | 82.0499 | 4.6573 | 88.0875
Irbesartan | 4000 | 4063.5075 | 375.994 | 9.2529 | 101.5877 | Irbesartan | KARL_1 | 400 | 398.5354 | 18.1485 | 4.5538 | 99.6338
Metoprolol | 400 | 402.5715 | 25.7392 | 6.3937 | 100.6429 | Metoprolol | KARL_1 | 4000 | 3846.5875 | 193.5942 | 5.0329 | 96.1647
Metoprolol | 2000 | 1850.9844 | 104.1979 | 5.6293 | 92.5492 | Metoprolol | KARL_1 | 2000 | 1803.6985 | 79.7405 | 4.4209 | 90.1849
Metoprolol | 4000 | 3881.0954 | 162.7733 | 4.194 | 97.0274 | Metoprolol | KARL_1 | 400 | 386.6391 | 5.3449 | 1.3824 | 96.6598
Propranolol | 400 | 425.8883 | 40.0702 | 9.4086 | 106.4721 | Propranolol | KARL_1 | 400 | 389.5706 | 41.6033 | 10.6793 | 97.3926
Propranolol | 2000 | 1897.856 | 217.5387 | 11.4623 | 94.8928 | Propranolol | KARL_1 | 2000 | 1694.3515 | 109.2945 | 6.4505 | 84.7176
Propranolol | 4000 | 3954.6389 | 96.0826 | 2.4296 | 98.866 | Propranolol | KARL_1 | 4000 | 3884.4115 | 114.8159 | 2.9558 | 97.1103
Tramadol | 400 | 403.8267 | 22.745 | 5.6324 | 100.9567 | Tramadol | KARL_1 | 2000 | 1816.0465 | 217.9923 | 12.0037 | 90.8023
Tramadol | 2000 | 1868.3779 | 134.3979 | 7.1933 | 93.4189 | Tramadol | KARL_1 | 4000 | 3683.2086 | 143.0951 | 3.8851 | 92.0802
Tramadol | 4000 | 3840.0734 | 211.7971 | 5.5154 | 96.0018 | Tramadol | KARL_1 | 400 | 379.2002 | 11.6354 | 3.0684 | 94.8
Valsartan | 400 | 443.7114 | 23.7173 | 5.3452 | 110.9278 | Valsartan | KARL_1 | 4000 | 3825.0799 | 348.4878 | 9.1106 | 95.627
Valsartan | 2000 | 1996.8451 | 135.1715 | 6.7693 | 99.8423 | Valsartan | KARL_1 | 2000 | 1987.3034 | 130.0974 | 6.5464 | 99.3652
Valsartan | 4000 | 3969.8389 | 421.8128 | 10.6254 | 99.246 | Valsartan | KARL_1 | 400 | 454.4746 | 9.727 | 2.1403 | 113.6186
Venlafaxine | 400 | 413.1256 | 20.7607 | 5.0253 | 103.2814 | Venlafaxine | KARL_1 | 400 | 396.9225 | 16.0665 | 4.0478 | 99.2306
Venlafaxine | 2000 | 1879.7585 | 133.7489 | 7.1152 | 93.9879 | Venlafaxine | KARL_1 | 2000 | 1766.3138 | 34.4643 | 1.9512 | 88.3157
Venlafaxine | 4000 | 3998.1182 | 164.4534 | 4.1133 | 99.953 | Venlafaxine | KARL_1 | 4000 | 3835.4315 | 70.8686 | 1.8477 | 95.8858
[6 empty] | Amisulpride | KARL_2 | 2000 | 2029.5286 | 118.9881 | 5.8628 | 101.4764
SI-6: specificity | [5 empty] | Amisulpride | KARL_2 | 4000 | 4258.9711 | 80.1062 | 1.8809 | 106.4743
Compound Name | Mean Peak Area in Blanks (n = 5) | Peak Area at LOQ | Blank/LOQ (%) |  |  | Amisulpride | KARL_2 | 400 | 423.8841 | 3.5831 | 0.8453 | 105.971
Amisulpride | 812 | 29603.7 | 2.7 |  |  | Candesartan | KARL_2 | 2000 | 1880.3274 | 142.348 | 7.5704 | 94.0164
Candesartan | 846.8 | 6377.6 | 13.3 |  |  | Candesartan | KARL_2 | 4000 | 3698.739 | 263.7632 | 7.1312 | 92.4685
Carbamazepin | 7896.4 | 65080.1 | 12.1 |  |  | Candesartan | KARL_2 | 400 | 424.9655 | 18.5147 | 4.3568 | 106.2414
Citalopram | 491.8 | 9191.6 | 5.4 |  |  | Carbamazepin | KARL_2 | 2000 | 1920.3764 | 86.6963 | 4.5145 | 96.0188
Clarithromycin | 0 | 21371.9 | 0 |  |  | Carbamazepin | KARL_2 | 400 | 430.5045 | 15.1317 | 3.5149 | 107.6261
Diclofenac | 81.5 | 3259.8 | 2.5 |  |  | Carbamazepin | KARL_2 | 4000 | 4028.9726 | 44.6463 | 1.1081 | 100.7243
Irbesartan | 0 | 5571.5 | 0 |  |  | Citalopram | KARL_2 | 400 | 429.2492 | 33.318 | 7.7619 | 107.3123
Metoprolol | 0 | 42233.8 | 0 |  |  | Citalopram | KARL_2 | 2000 | 1909.2142 | 57.8825 | 3.0317 | 95.4607
Propranolol | 12222.9 | 65478.8 | 18.7 |  |  | Citalopram | KARL_2 | 4000 | 4216.8951 | 90.4963 | 2.146 | 105.4224
Tramadol | 0 | 44012.5 | 0 |  |  | Clarithromycin | KARL_2 | 4000 | 4095.5295 | 217.4022 | 5.3083 | 102.3882
Valsartan | 234.4 | 7318.4 | 3.2 |  |  | Clarithromycin | KARL_2 | 2000 | 1803.3891 | 62.4856 | 3.4649 | 90.1695
Venlafaxine | 0 | 22426.7 | 0 |  |  | Clarithromycin | KARL_2 | 400 | 460.9879 | 0.45722 | 0.099182 | 115.247
[6 empty] | Diclofenac | KARL_2 | 10000 | 10144.2876 | 335.0993 | 3.3033 | 101.4429
[6 empty] | Diclofenac | KARL_2 | 2000 | 2094.5445 | 57.5498 | 2.7476 | 104.7272
[6 empty] | Diclofenac | KARL_2 | 20000 | 19910.8164 | 104.8944 | 0.52682 | 99.5541
[6 empty] | Irbesartan | KARL_2 | 2000 | 2052.5822 | 247.7901 | 12.0721 | 102.6291
[6 empty] | Irbesartan | KARL_2 | 4000 | 4196.6958 | 266.0479 | 6.3395 | 104.9174
[6 empty] | Irbesartan | KARL_2 | 400 | 457.6482 | 11.0683 | 2.4185 | 114.412
[6 empty] | Metoprolol | KARL_2 | 2000 | 1904.4229 | 134.7822 | 7.0773 | 95.2211
[6 empty] | Metoprolol | KARL_2 | 400 | 399.5726 | 22.6198 | 5.661 | 99.8932
[6 empty] | Metoprolol | KARL_2 | 4000 | 3957.8311 | 140.6046 | 3.5526 | 98.9458
[6 empty] | Propranolol | KARL_2 | 2000 | 2004.4714 | 169.4274 | 8.4525 | 100.2236
[6 empty] | Propranolol | KARL_2 | 400 | 443.6435 | 25.3141 | 5.7059 | 110.9109
[6 empty] | Propranolol | KARL_2 | 4000 | 4024.2284 | 69.8334 | 1.7353 | 100.6057
[6 empty] | Tramadol | KARL_2 | 2000 | 1916.7767 | 81.9536 | 4.2756 | 95.8388
[6 empty] | Tramadol | KARL_2 | 4000 | 3813.1837 | 127.3084 | 3.3386 | 95.3296
[6 empty] | Tramadol | KARL_2 | 400 | 404.8944 | 11.4192 | 2.8203 | 101.2236
[6 empty] | Valsartan | KARL_2 | 2000 | 2076.4301 | 160.1469 | 7.7126 | 103.8215
[6 empty] | Valsartan | KARL_2 | 4000 | 3716.1484 | 122.5497 | 3.2978 | 92.9037
[6 empty] | Valsartan | KARL_2 | 400 | 457.2337 | 5.9016 | 1.2907 | 114.3084
[6 empty] | Venlafaxine | KARL_2 | 2000 | 1960.6159 | 147.4478 | 7.5205 | 98.0308
[6 empty] | Venlafaxine | KARL_2 | 400 | 405.5093 | 6.1623 | 1.5196 | 101.3773
[6 empty] | Venlafaxine | KARL_2 | 4000 | 3984.6332 | 41.7506 | 1.0478 | 99.6158
[6 empty] | Amisulpride | KARL_3 | 400 | 437.9644 | 26.9041 | 6.143 | 109.4911
[6 empty] | Amisulpride | KARL_3 | 4000 | 4126.2395 | 104.0676 | 2.5221 | 103.156
[6 empty] | Amisulpride | KARL_3 | 2000 | 1933.1508 | 25.2537 | 1.3063 | 96.6575
[6 empty] | Candesartan | KARL_3 | 400 | 436.6793 | 28.0636 | 6.4266 | 109.1698
[6 empty] | Candesartan | KARL_3 | 4000 | 3986.0945 | 107.705 | 2.702 | 99.6524
[6 empty] | Candesartan | KARL_3 | 2000 | 1913.3781 | 50.4681 | 2.6376 | 95.6689
[6 empty] | Carbamazepin | KARL_3 | 4000 | 3670.9221 | 194.6718 | 5.3031 | 91.7731
[6 empty] | Carbamazepin | KARL_3 | 400 | 425.3496 | 19.5617 | 4.599 | 106.3374
[6 empty] | Carbamazepin | KARL_3 | 2000 | 1809.7498 | 30.3931 | 1.6794 | 90.4875
[6 empty] | Citalopram | KARL_3 | 4000 | 4437.6004 | 345.9344 | 7.7955 | 110.94
[6 empty] | Citalopram | KARL_3 | 2000 | 2011.9455 | 125.4561 | 6.2356 | 100.5973
[6 empty] | Citalopram | KARL_3 | 400 | 454.7325 | 17.1799 | 3.778 | 113.6831
[6 empty] | Clarithromycin | KARL_3 | 400 | 454.9313 | 50.1795 | 11.0301 | 113.7328
[6 empty] | Clarithromycin | KARL_3 | 4000 | 4208.8657 | 381.569 | 9.0658 | 105.2216
[6 empty] | Clarithromycin | KARL_3 | 2000 | 1920.1457 | 108.58 | 5.6548 | 96.0073
[6 empty] | Diclofenac | KARL_3 | 10000 | 9454.4562 | 271.8549 | 2.8754 | 94.5446
[6 empty] | Diclofenac | KARL_3 | 20000 | 20954.0114 | 581.3202 | 2.7743 | 104.7701
[6 empty] | Diclofenac | KARL_3 | 2000 | 2120.161 | 42.004 | 1.9812 | 106.0081
[6 empty] | Irbesartan | KARL_3 | 400 | 447.5149 | 21.9005 | 4.8938 | 111.8787
[6 empty] | Irbesartan | KARL_3 | 4000 | 4316.3814 | 127.367 | 2.9508 | 107.9095
[6 empty] | Irbesartan | KARL_3 | 2000 | 2099.7397 | 11.6364 | 0.55418 | 104.987
[6 empty] | Metoprolol | KARL_3 | 400 | 421.5028 | 34.3213 | 8.1426 | 105.3757
[6 empty] | Metoprolol | KARL_3 | 2000 | 1841.7556 | 107.2641 | 5.824 | 92.0878
[6 empty] | Metoprolol | KARL_3 | 4000 | 3838.8677 | 188.2517 | 4.9038 | 95.9717
[6 empty] | Propranolol | KARL_3 | 2000 | 2043.1898 | 222.2007 | 10.8752 | 102.1595
[6 empty] | Propranolol | KARL_3 | 4000 | 3955.2768 | 64.8505 | 1.6396 | 98.8819
[6 empty] | Propranolol | KARL_3 | 400 | 453.7321 | 6.5458 | 1.4426 | 113.433
[6 empty] | Tramadol | KARL_3 | 4000 | 4023.828 | 232.6056 | 5.7807 | 100.5957
[6 empty] | Tramadol | KARL_3 | 2000 | 1874.2766 | 51.5064 | 2.7481 | 93.7138
[6 empty] | Tramadol | KARL_3 | 400 | 427.3855 | 7.7219 | 1.8068 | 106.8464
[6 empty] | Valsartan | KARL_3 | 400 | 407.2831 | 11.3378 | 2.7838 | 101.8208
[6 empty] | Valsartan | KARL_3 | 2000 | 1891.7803 | 37.4567 | 1.98 | 94.589
[6 empty] | Valsartan | KARL_3 | 4000 | 4567.5134 | 46.7602 | 1.0238 | 114.1878
[6 empty] | Venlafaxine | KARL_3 | 2000 | 1928.6392 | 127.916 | 6.6324 | 96.432
[6 empty] | Venlafaxine | KARL_3 | 4000 | 4174.2899 | 122.0096 | 2.9229 | 104.3572
[6 empty] | Venlafaxine | KARL_3 | 400 | 436.9449 | 9.8007 | 2.243 | 109.2362

## recovery and matrixeffects
SI-7: matrix effects | [5 empty] | SI-8: recovery
standards | [5 empty] | standards
compound | sample_level | theoretical_amount | peak_area_C | peak_area_B | matrixeffect_percent | compound | sample_level | theoretical_amount | peak_area_A | peak_area_B | percent | recovery
Amisulpride | Level 1 | 400 | 12368185.7646 | 6571905.7339 | 53.1356 | Amisulpride | Level 1 | 400 | 7291146.8281 | 6571905.7339 | 110.9442 | 10.9442
Amisulpride | Level 2 | 2000 | 57575575.8883 | 32740718.181 | 56.8656 | Amisulpride | Level 2 | 2000 | 35921380.0414 | 32740718.181 | 109.7147 | 9.7147
Amisulpride | Level 3 | 4000 | 121353660.0543 | 66681117.8784 | 54.9478 | Amisulpride | Level 3 | 4000 | 73087883.9409 | 66681117.8784 | 109.6081 | 9.6081
Candesartan | Level 1 | 400 | 460738.6777 | 611803.7423 | 132.7876 | Candesartan | Level 1 | 400 | 498078.9788 | 611803.7423 | 81.4116 | -18.5884
Candesartan | Level 2 | 2000 | 2157030.8401 | 2709332.1283 | 125.6047 | Candesartan | Level 2 | 2000 | 2374429.9961 | 2709332.1283 | 87.6389 | -12.3611
Candesartan | Level 3 | 4000 | 4681878.9435 | 5381157.3671 | 114.9359 | Candesartan | Level 3 | 4000 | 4632723.0361 | 5381157.3671 | 86.0916 | -13.9084
Carbamazepine | Level 1 | 400 | 5997636.6311 | 5876989.1105 | 97.9884 | Carbamazepine | Level 1 | 400 | 6431135.1602 | 5876989.1105 | 109.4291 | 9.4291
Carbamazepine | Level 2 | 2000 | 26493928.052 | 27583028.6588 | 104.1108 | Carbamazepine | Level 2 | 2000 | 29608062.7271 | 27583028.6588 | 107.3416 | 7.3416
Carbamazepine | Level 3 | 4000 | 51774372.6099 | 54794312.2468 | 105.8329 | Carbamazepine | Level 3 | 4000 | 57558547.1257 | 54794312.2468 | 105.0447 | 5.0447
Citalopram | Level 1 | 400 | 19417386.3745 | 11866292.8422 | 61.1117 | Citalopram | Level 1 | 400 | 13269333.8495 | 11866292.8422 | 111.8238 | 11.8238
Citalopram | Level 2 | 2000 | 94725083.4762 | 59157048.9165 | 62.4513 | Citalopram | Level 2 | 2000 | 67003610.7155 | 59157048.9165 | 113.264 | 13.264
Citalopram | Level 3 | 4000 | 191532211.593 | 124803790.0217 | 65.1607 | Citalopram | Level 3 | 4000 | 137830600.3789 | 124803790.0217 | 110.4378 | 10.4378
Clarithromycin | Level 1 | 400 | 11179615.0328 | 7044858.4447 | 63.0152 | Clarithromycin | Level 1 | 400 | 6853002.0515 | 7044858.4447 | 97.2766 | -2.7234
Clarithromycin | Level 2 | 2000 | 50599572.6118 | 33518099.5966 | 66.2419 | Clarithromycin | Level 2 | 2000 | 34209542.4271 | 33518099.5966 | 102.0629 | 2.0629
Clarithromycin | Level 3 | 4000 | 110883111.3622 | 70319206.8696 | 63.4174 | Clarithromycin | Level 3 | 4000 | 68518265.8464 | 70319206.8696 | 97.4389 | -2.5611
diclofenac | Level 1 | 2000 | 535812.2214 | 596139.2955 | 111.259 | diclofenac | Level 1 | 2000 | 507655.8321 | 596139.2955 | 85.1573 | -14.8427
diclofenac | Level 2 | 10000 | 2550332.163 | 2766694.3581 | 108.4837 | diclofenac | Level 2 | 10000 | 2417113.8689 | 2766694.3581 | 87.3647 | -12.6353
diclofenac | Level 3 | 20000 | 5606071.4773 | 5656216.1538 | 100.8945 | diclofenac | Level 3 | 20000 | 4609245.0819 | 5656216.1538 | 81.4899 | -18.5101
Irbesartan | Level 1 | 400 | 4937224.7786 | 8705766.0276 | 176.3291 | Irbesartan | Level 1 | 400 | 6443180.4217 | 8705766.0276 | 74.0105 | -25.9895
Irbesartan | Level 2 | 2000 | 23275644.8677 | 33139721.3332 | 142.3794 | Irbesartan | Level 2 | 2000 | 31003806.0458 | 33139721.3332 | 93.5548 | -6.4452
Irbesartan | Level 3 | 4000 | 51071576.5757 | 62267344.6653 | 121.9217 | Irbesartan | Level 3 | 4000 | 56738716.2513 | 62267344.6653 | 91.1211 | -8.8789
Metoprolol | Level 1 | 400 | 6686751.3198 | 7768235.9966 | 116.1735 | Metoprolol | Level 1 | 400 | 8712049.1659 | 7768235.9966 | 112.1496 | 12.1496
Metoprolol | Level 2 | 2000 | 31769404.8262 | 38310527.3958 | 120.5894 | Metoprolol | Level 2 | 2000 | 43630864.5887 | 38310527.3958 | 113.8874 | 13.8874
Metoprolol | Level 3 | 4000 | 64738240.5582 | 78247742.7208 | 120.8679 | Metoprolol | Level 3 | 4000 | 88790750.6514 | 78247742.7208 | 113.4739 | 13.4739
Propranolol | Level 1 | 400 | 7096977.8049 | 6855063.7932 | 96.5913 | Propranolol | Level 1 | 400 | 7138984.506 | 6855063.7932 | 104.1418 | 4.1418
Propranolol | Level 2 | 2000 | 34888616.387 | 29542413.5987 | 84.6764 | Propranolol | Level 2 | 2000 | 34948397.7072 | 29542413.5987 | 118.2991 | 18.2991
Propranolol | Level 3 | 4000 | 67775950.6759 | 62880972.4486 | 92.7777 | Propranolol | Level 3 | 4000 | 71575313.7779 | 62880972.4486 | 113.8267 | 13.8267
Tramadol | Level 1 | 400 | 5674336.4626 | 4305906.7221 | 75.8839 | Tramadol | Level 1 | 400 | 4479334.0113 | 4305906.7221 | 104.0277 | 4.0277
Tramadol | Level 2 | 2000 | 27141730.5969 | 19866040.9042 | 73.1937 | Tramadol | Level 2 | 2000 | 21537530.2922 | 19866040.9042 | 108.4138 | 8.4138
Tramadol | Level 3 | 4000 | 58415725.9473 | 39569995.8167 | 67.7386 | Tramadol | Level 3 | 4000 | 42957285.1171 | 39569995.8167 | 108.5602 | 8.5602
Valsartan | Level 1 | 400 | 217481.482 | 378300.1154 | 173.9459 | Valsartan | Level 1 | 400 | 206525.649 | 378300.1154 | 54.5931 | -45.4069
Valsartan | Level 2 | 2000 | 1035372.2285 | 1351625.5798 | 130.5449 | Valsartan | Level 2 | 2000 | 1025083.1866 | 1351625.5798 | 75.8408 | -24.1592
Valsartan | Level 3 | 4000 | 2267742.0926 | 2612818.7318 | 115.2167 | Valsartan | Level 3 | 4000 | 1832727.0718 | 2612818.7318 | 70.1437 | -29.8563
Venlafaxine | Level 1 | 400 | 7438603.9569 | 5515235.599 | 74.1434 | Venlafaxine | Level 1 | 400 | 5681055.0393 | 5515235.599 | 103.0066 | 3.0066
Venlafaxine | Level 2 | 2000 | 35582152.3917 | 25465413.2735 | 71.5679 | Venlafaxine | Level 2 | 2000 | 26385068.2277 | 25465413.2735 | 103.6114 | 3.6114
Venlafaxine | Level 3 | 4000 | 76216202.7664 | 51767119.7402 | 67.9214 | Venlafaxine | Level 3 | 4000 | 53784761.4937 | 51767119.7402 | 103.8975 | 3.8975
ISTD | [5 empty] | ISTD
compound | Column1 | sample_level | peak_area_C | peak_area_B | matrixeffect_percent | sample_level | theoretical_amount | compound | peak_area_A | peak_area_B | percent | recovery
Level 1 | 400 | Amisulpride-d5 | 32665215.4057 | 17913871.3258 | 54.8408 | Level 1 | 100000 | Amisulpride-d5 | 17683499.8172 | 17913871.3258 | 98.714 | -1.286
Level 2 | 2000 | Amisulpride-d5 | 31606533.4981 | 17497382.6301 | 55.36 | Level 2 | 100000 | Amisulpride-d5 | 18115712.9511 | 17497382.6301 | 103.5338 | 3.5338
Level 3 | 4000 | Amisulpride-d5 | 31602666.1912 | 18718451.9584 | 59.2306 | Level 3 | 100000 | Amisulpride-d5 | 19051327.1395 | 18718451.9584 | 101.7783 | 1.7783
Level 1 | 400 | Candesartan-d5 | 18897848.2175 | 23603150.5032 | 124.8986 | Level 1 | 100000 | Candesartan-d5 | 18264571.2711 | 23603150.5032 | 77.3819 | -22.6181
Level 2 | 2000 | Candesartan-d5 | 17905560.331 | 20889051.6753 | 116.6624 | Level 2 | 100000 | Candesartan-d5 | 17815902.4351 | 20889051.6753 | 85.2882 | -14.7118
Level 3 | 4000 | Candesartan-d5 | 17387233.9086 | 21192597.206 | 121.886 | Level 3 | 100000 | Candesartan-d5 | 17550673.6333 | 21192597.206 | 82.8151 | -17.1849
Level 1 | 400 | Carbamazepine-d10 | 26224917.9905 | 24001242.6755 | 91.5208 | Level 1 | 100000 | Carbamazepine-d10 | 24958972.5981 | 24001242.6755 | 103.9903 | 3.9903
Level 2 | 2000 | Carbamazepine-d10 | 24588232.5804 | 23016963.9714 | 93.6097 | Level 2 | 100000 | Carbamazepine-d10 | 23647856.4088 | 23016963.9714 | 102.741 | 2.741
Level 3 | 4000 | Carbamazepine-d10 | 23437429.558 | 24207259.7787 | 103.2846 | Level 3 | 100000 | Carbamazepine-d10 | 23559744.9133 | 24207259.7787 | 97.3251 | -2.6749
Level 1 | 400 | Citalopram-d6 | 50785997.3282 | 31391496.5973 | 61.8113 | Level 1 | 100000 | Citalopram-d6 | 33505674.2874 | 31391496.5973 | 106.7349 | 6.7349
Level 2 | 2000 | Citalopram-d6 | 49477694.5989 | 30683121.3927 | 62.014 | Level 2 | 100000 | Citalopram-d6 | 32695263.2075 | 30683121.3927 | 106.5578 | 6.5578
Level 3 | 4000 | Citalopram-d6 | 48761469.2534 | 32702780.4708 | 67.0668 | Level 3 | 100000 | Citalopram-d6 | 34198853.7345 | 32702780.4708 | 104.5748 | 4.5748
Level 1 | 400 | Clarithromycin-d3 | 9881462.0873 | 6713767.1892 | 67.9431 | Level 1 | 100000 | Clarithromycin-d3 | 6703460.2575 | 6713767.1892 | 99.8465 | -0.15352
Level 2 | 2000 | Clarithromycin-d3 | 9687464.8837 | 6371570.8534 | 65.7713 | Level 2 | 100000 | Clarithromycin-d3 | 6608833.985 | 6371570.8534 | 103.7238 | 3.7238
Level 3 | 4000 | Clarithromycin-d3 | 9771467.8303 | 7116120.4799 | 72.8255 | Level 3 | 100000 | Clarithromycin-d3 | 6835248.2259 | 7116120.4799 | 96.053 | -3.947
Level 1 | 400 | Irbesartan-d4 | 13195151.3939 | 20425699.3623 | 154.797 | Level 1 | 100000 | Irbesartan-d4 | 14895567.2409 | 20425699.3623 | 72.9256 | -27.0744
Level 2 | 2000 | Irbesartan-d4 | 12419404.3614 | 15765024.3998 | 126.9387 | Level 2 | 100000 | Irbesartan-d4 | 14134068.6924 | 15765024.3998 | 89.6546 | -10.3454
Level 3 | 4000 | Irbesartan-d4 | 12627745.9901 | 15853481.7844 | 125.5448 | Level 3 | 100000 | Irbesartan-d4 | 13291045.0824 | 15853481.7844 | 83.8368 | -16.1632
Level 1 | 400 | Metoprolol-d7 | 21093779.9754 | 23432467.1943 | 111.0871 | Level 1 | 100000 | Metoprolol-d7 | 24624970.3869 | 23432467.1943 | 105.0891 | 5.0891
Level 2 | 2000 | Metoprolol-d7 | 20234206.0264 | 22800185.8898 | 112.6814 | Level 2 | 100000 | Metoprolol-d7 | 24386259.9008 | 22800185.8898 | 106.9564 | 6.9564
Level 3 | 4000 | Metoprolol-d7 | 19916388.092 | 23901572.2518 | 120.0096 | Level 3 | 100000 | Metoprolol-d7 | 24871926.6746 | 23901572.2518 | 104.0598 | 4.0598
Level 1 | 400 | Propranolol-d7 | 24025337.9834 | 21025215.9947 | 87.5127 | Level 1 | 100000 | Propranolol-d7 | 23069717.2766 | 21025215.9947 | 109.724 | 9.724
Level 2 | 2000 | Propranolol-d7 | 23532485.0375 | 21238207.6238 | 90.2506 | Level 2 | 100000 | Propranolol-d7 | 23181243.2163 | 21238207.6238 | 109.1488 | 9.1488
Level 3 | 4000 | Propranolol-d7 | 22728322.806 | 22723951.3077 | 99.9808 | Level 3 | 100000 | Propranolol-d7 | 23002254.9453 | 22723951.3077 | 101.2247 | 1.2247
Level 1 | 400 | Tramadol-d3 | 56941463.0939 | 38850554.7845 | 68.2289 | Level 1 | 100000 | Tramadol-d3 | 38433424.6854 | 38850554.7845 | 98.9263 | -1.0737
Level 2 | 2000 | Tramadol-d3 | 56054314.7469 | 37484655.1401 | 66.872 | Level 2 | 100000 | Tramadol-d3 | 38121010.0447 | 37484655.1401 | 101.6976 | 1.6976
Level 3 | 4000 | Tramadol-d3 | 54919705.1642 | 38299420.5203 | 69.7371 | Level 3 | 100000 | Tramadol-d3 | 39040889.9487 | 38299420.5203 | 101.936 | 1.936
Level 1 | 400 | Valsartan-d9 | 2052001.5588 | 3980281.075 | 193.9707 | Level 1 | 100000 | Valsartan-d9 | 1762537.5437 | 3980281.075 | 44.2817 | -55.7183
Level 2 | 2000 | Valsartan-d9 | 2018193.3335 | 2785175.4613 | 138.0034 | Level 2 | 100000 | Valsartan-d9 | 1767597.6023 | 2785175.4613 | 63.4645 | -36.5355
Level 3 | 4000 | Valsartan-d9 | 2020025.612 | 2520595.7811 | 124.7804 | Level 3 | 100000 | Valsartan-d9 | 1718417.5748 | 2520595.7811 | 68.1751 | -31.8249
Level 1 | 400 | Venlafaxine-d6 | 46176089.079 | 31881816.7401 | 69.044 | Level 1 | 100000 | Venlafaxine-d6 | 30118672.6578 | 31881816.7401 | 94.4698 | -5.5302
Level 2 | 2000 | Venlafaxine-d6 | 45635560.455 | 29524797.4422 | 64.6969 | Level 2 | 100000 | Venlafaxine-d6 | 30285252.3319 | 29524797.4422 | 102.5756 | 2.5756
Level 3 | 4000 | Venlafaxine-d6 | 42308978.4619 | 30537134.139 | 72.1765 | Level 3 | 100000 | Venlafaxine-d6 | 29916277.6815 | 30537134.139 | 97.9669 | -2.0331

## storage tests
SI-9: QC storage at 25 degree celsius for 24 h
sample_level | compound | conc_1 | n | conc_2 | recovery_percent
Level 1 | Amisulpride | 370.9744 | 3 | 373.4291 | 100.6617
Level 2 | Amisulpride | 1873.6676 | 3 | 1723.3541 | 91.9776
Level 3 | Amisulpride | 3876.5891 | 3 | 3364.1315 | 86.7807
Level 1 | Candesartan | 380.3474 | 3 | 378.1447 | 99.4209
Level 2 | Candesartan | 1915.9279 | 3 | 1831.7973 | 95.6089
Level 3 | Candesartan | 4080.7314 | 3 | 3741.7911 | 91.6941
Level 1 | Carbamazepine | 404.3182 | 3 | 391.4977 | 96.8291
Level 2 | Carbamazepine | 2044.4464 | 3 | 1884.1879 | 92.1613
Level 3 | Carbamazepine | 4061.5966 | 3 | 3783.6707 | 93.1572
Level 1 | Citalopram | 407.9609 | 3 | 398.6747 | 97.7238
Level 2 | Citalopram | 1962.7337 | 3 | 1805.0407 | 91.9656
Level 3 | Citalopram | 3950.6467 | 3 | 3906.5406 | 98.8836
Level 1 | Clarithromycin | 384.5668 | 3 | 333.4901 | 86.7184
Level 2 | Clarithromycin | 2002.7006 | 3 | 1681.091 | 83.9412
Level 3 | Clarithromycin | 4179.4129 | 3 | 3541.475 | 84.7362
Level 1 | Diclofenac | 1985.9349 | 3 | 1907.6272 | 96.0569
Level 2 | Diclofenac | 9520.8344 | 3 | 8405.9782 | 88.2904
Level 3 | Diclofenac | 17791.7365 | 3 | 18090.5476 | 101.6795
Level 1 | Irbesartan | 400.9606 | 3 | 397.0496 | 99.0246
Level 2 | Irbesartan | 2021.1532 | 3 | 1839.4397 | 91.0094
Level 3 | Irbesartan | 4090.5403 | 3 | 3786.0562 | 92.5564
Level 1 | Metoprolol | 418.9927 | 3 | 412.1987 | 98.3785
Level 2 | Metoprolol | 1928.5484 | 3 | 1829.6082 | 94.8697
Level 3 | Metoprolol | 3961.8767 | 3 | 3802.818 | 95.9853
Level 1 | Propranolol | 386.6294 | 3 | 389.6563 | 100.7829
Level 2 | Propranolol | 1989.1051 | 3 | 1900.6518 | 95.5531
Level 3 | Propranolol | 3962.2586 | 3 | 3827.3651 | 96.5955
Level 1 | Tramadol | 385.7001 | 3 | 401.4635 | 104.0869
Level 2 | Tramadol | 1969.1841 | 3 | 1898.2753 | 96.3991
Level 3 | Tramadol | 4211.503 | 3 | 3790.7497 | 90.0094
Level 1 | Valsartan | 412.0683 | 3 | 364.3715 | 88.425
Level 2 | Valsartan | 1961.3003 | 3 | 1711.1683 | 87.2466
Level 3 | Valsartan | 4096.0629 | 3 | 3446.3033 | 84.137
Level 1 | Venlafaxine | 391.6851 | 3 | 404.3561 | 103.235
Level 2 | Venlafaxine | 1969.7208 | 3 | 1902.3845 | 96.5814
Level 3 | Venlafaxine | 4174.9556 | 3 | 3721.2603 | 89.1329
SI-10: QC storage at 4 degree celsius for 14 days
sample_level | compound | conc_1 | conc_2 | recovery_percent | n
Level 1 | Amisulpride | 370.9744 | 413.3116 | 111.4124 | 3
Level 2 | Amisulpride | 1873.6676 | 1883.4559 | 100.5224 | 3
Level 3 | Amisulpride | 3876.5891 | 3728.5109 | 96.1802 | 3
Level 1 | Candesartan | 380.3474 | 398.3203 | 104.7254 | 3
Level 2 | Candesartan | 1915.9279 | 1876.1553 | 97.9241 | 3
Level 3 | Candesartan | 4080.7314 | 3696.78 | 90.5911 | 3
Level 1 | Carbamazepine | 404.3182 | 416.3677 | 102.9802 | 3
Level 2 | Carbamazepine | 2044.4464 | 1865.6122 | 91.2527 | 3
Level 3 | Carbamazepine | 4061.5966 | 3792.3132 | 93.37 | 3
Level 1 | Citalopram | 407.9609 | 471.6491 | 115.6113 | 3
Level 2 | Citalopram | 1962.7337 | 1970.213 | 100.3811 | 3
Level 3 | Citalopram | 3950.6467 | 4073.2304 | 103.1029 | 3
Level 1 | Clarithromycin | 384.5668 | 429.8775 | 111.7823 | 3
Level 2 | Clarithromycin | 2002.7006 | 2116.1042 | 105.6625 | 3
Level 3 | Clarithromycin | 4179.4129 | 4252.8922 | 101.7581 | 3
Level 1 | Diclofenac | 1985.9349 | 2013.8991 | 101.4081 | 3
Level 2 | Diclofenac | 9520.8344 | 10608.0344 | 111.4192 | 3
Level 3 | Diclofenac | 17791.7365 | 19729.8629 | 110.8934 | 3
Level 1 | Irbesartan | 400.9606 | 410.5523 | 102.3922 | 3
Level 2 | Irbesartan | 2021.1532 | 1857.768 | 91.9162 | 3
Level 3 | Irbesartan | 4090.5403 | 3760.3504 | 91.928 | 3
Level 1 | Metoprolol | 418.9927 | 401.6359 | 95.8575 | 3
Level 2 | Metoprolol | 1928.5484 | 1836.6703 | 95.2359 | 3
Level 3 | Metoprolol | 3961.8767 | 3791.0955 | 95.6894 | 3
Level 1 | Propranolol | 386.6294 | 413.2613 | 106.8882 | 2
Level 2 | Propranolol | 1989.1051 | 1964.3434 | 98.7551 | 3
Level 3 | Propranolol | 3962.2586 | 3882.2773 | 97.9814 | 3
Level 1 | Tramadol | 385.7001 | 412.495 | 106.9471 | 3
Level 2 | Tramadol | 1969.1841 | 1890.8449 | 96.0217 | 3
Level 3 | Tramadol | 4211.503 | 3906.163 | 92.7499 | 3
Level 1 | Valsartan | 412.0683 | 430.2238 | 104.406 | 3
Level 2 | Valsartan | 1961.3003 | 1821.7544 | 92.885 | 3
Level 3 | Valsartan | 4096.0629 | 4226.2537 | 103.1784 | 3
Level 1 | Venlafaxine | 391.6851 | 423.503 | 108.1233 | 3
Level 2 | Venlafaxine | 1969.7208 | 1920.0094 | 97.4762 | 3
Level 3 | Venlafaxine | 4174.9556 | 3821.1834 | 91.5263 | 3
SI11: QC storage at 10 degree celsius for 24h in autosampler
sample_level | compound | conc_1 | n | conc_2 | recovery_percent
Level 1 | Amisulpride | 381.0807 | 3 | 376.0218 | 98.6725
Level 2 | Amisulpride | 1910.1806 | 2 | 1894.7775 | 99.1936
Level 3 | Amisulpride | 3965.4611 | 3 | 3979.4486 | 100.3527
Level 1 | Candesartan | 376.8002 | 3 | 363.5218 | 96.476
Level 2 | Candesartan | 1793.9242 | 2 | 1725.5869 | 96.1906
Level 3 | Candesartan | 3679.8881 | 3 | 3734.3583 | 101.4802
Level 1 | Carbamazepine | 419.2154 | 3 | 410.8143 | 97.996
Level 2 | Carbamazepine | 2114.4391 | 2 | 2209.9819 | 104.5186
Level 3 | Carbamazepine | 3965.9208 | 3 | 3884.0783 | 97.9364
Level 1 | Citalopram | 368.646 | 3 | 379.5411 | 102.9554
Level 2 | Citalopram | 2034.4769 | 2 | 2019.8474 | 99.2809
Level 3 | Citalopram | 3907.6197 | 3 | 3733.7691 | 95.551
Level 1 | Clarithromycin | 330.8978 | 3 | 309.2749 | 93.4654
Level 2 | Clarithromycin | 1544.7818 | 2 | 1574.4996 | 101.9238
Level 3 | Clarithromycin | 3332.4715 | 3 | 3399.971 | 102.0255
Level 1 | Diclofenac | 3944.2532 | 3 | 3892.2305 | 98.6811
Level 2 | Diclofenac | 13651.5172 | 2 | 15046.1483 | 110.2159
Level 3 | Diclofenac | 35879.2817 | 3 | 40235.8304 | 112.1422
Level 1 | Irbesartan | 442.501 | 3 | 440.5607 | 99.5615
Level 2 | Irbesartan | 2068.0419 | 2 | 2019.8459 | 97.6695
Level 3 | Irbesartan | 3923.8059 | 3 | 3912.6891 | 99.7167
Level 1 | Metoprolol | 701.2478 | 3 | 707.2858 | 100.861
Level 2 | Metoprolol | 2846.257 | 2 | 2866.5806 | 100.714
Level 3 | Metoprolol | 7097.3703 | 3 | 7153.2445 | 100.7873
Level 1 | Propranolol | 503.8182 | 3 | 494.8963 | 98.2291
Level 2 | Propranolol | 2269.7101 | 2 | 2285.4124 | 100.6918
Level 3 | Propranolol | 4688.3927 | 3 | 4809.4304 | 102.5816
Level 1 | Tramadol | 442.4866 | 3 | 421.4869 | 95.2542
Level 2 | Tramadol | 2066.2928 | 2 | 2047.0482 | 99.0686
Level 3 | Tramadol | 4175.1609 | 3 | 4378.2031 | 104.8631
Level 1 | Valsartan | 484.652 | 3 | 433.9576 | 89.54
Level 2 | Valsartan | 1705.0937 | 2 | 1758.7691 | 103.1479
Level 3 | Valsartan | 3829.2653 | 3 | 4027.6198 | 105.18
Level 1 | Venlafaxine | 504.7317 | 3 | 484.8233 | 96.0556
Level 2 | Venlafaxine | 2459.2578 | 2 | 2390.1357 | 97.1893
Level 3 | Venlafaxine | 4984.6714 | 3 | 5227.5787 | 104.8731
SI12: QC storage in 3 freeze and thaw cycles
sample_level | compound | conc_1 | n.y | conc_2 | recovery_percent
Level 1 | Amisulpride | 413.3116 | 3 | 473.7504 | 114.623
Level 2 | Amisulpride | 1883.4559 | 2 | 1852.3849 | 98.3503
Level 3 | Amisulpride | 3728.5109 | 2 | 3729.6023 | 100.0293
Level 1 | Candesartan | 398.3203 | 3 | 370.8719 | 93.109
Level 2 | Candesartan | 1876.1553 | 2 | 1838.99 | 98.0191
Level 3 | Candesartan | 3696.78 | 2 | 3623.0222 | 98.0048
Level 1 | Carbamazepine | 416.3677 | 3 | 406.6686 | 97.6706
Level 2 | Carbamazepine | 1865.6122 | 2 | 1904.5666 | 102.088
Level 3 | Carbamazepine | 3792.3132 | 2 | 3696.0824 | 97.4625
Level 1 | Citalopram | 471.6491 | 3 | 402.4179 | 85.3215
Level 2 | Citalopram | 1970.213 | 2 | 1969.9455 | 99.9864
Level 3 | Citalopram | 4073.2304 | 2 | 4050.2253 | 99.4352
Level 1 | Clarithromycin | 429.8775 | 3 | 410.627 | 95.5219
Level 2 | Clarithromycin | 2116.1042 | 2 | 2063.0159 | 97.4912
Level 3 | Clarithromycin | 4252.8922 | 2 | 3861.5217 | 90.7975
Level 1 | Diclofenac | 2013.8991 | 3 | 1382.9696 | 68.6712
Level 2 | Diclofenac | 10608.0344 | 2 | 10313.0871 | 97.2196
Level 3 | Diclofenac | 19729.8629 | 2 | 17675.2797 | 89.5864
Level 1 | Irbesartan | 410.5523 | 3 | 339.7715 | 82.7596
Level 2 | Irbesartan | 1857.768 | 2 | 1730.1444 | 93.1303
Level 3 | Irbesartan | 3760.3504 | 2 | 3491.7839 | 92.8579
Level 1 | Metoprolol | 401.6359 | 3 | 414.6267 | 103.2345
Level 2 | Metoprolol | 1836.6703 | 2 | 1864.7678 | 101.5298
Level 3 | Metoprolol | 3791.0955 | 2 | 3637.2099 | 95.9409
Level 1 | Propranolol | 413.2613 | 3 | 399.9323 | 96.7747
Level 2 | Propranolol | 1964.3434 | 2 | 1974.7841 | 100.5315
Level 3 | Propranolol | 3882.2773 | 2 | 3780.5928 | 97.3808
Level 1 | Tramadol | 412.495 | 3 | 367.2406 | 89.0291
Level 2 | Tramadol | 1890.8449 | 2 | 1781.6 | 94.2224
Level 3 | Tramadol | 3906.163 | 2 | 3845.0217 | 98.4347
Level 1 | Valsartan | 430.2238 | 3 | 368.1888 | 85.5807
Level 2 | Valsartan | 1821.7544 | 2 | 1639.2552 | 89.9822
Level 3 | Valsartan | 4226.2537 | 2 | 3521.7078 | 83.3293
Level 1 | Venlafaxine | 423.503 | 3 | 388.9111 | 91.832
Level 2 | Venlafaxine | 1920.0094 | 2 | 1891.5631 | 98.5184
Level 3 | Venlafaxine | 3821.1834 | 2 | 3614.7124 | 94.5967
SI13: working solution storage at 25 degree celsius for 24 h
sample_level | compound | conc_1 | n | conc_2 | recovery_percent
Level 4 | Amisulpride | 5016.3807 | 3 | 4428.8711 | 88.2882
Level 5 | Amisulpride | 484.4306 | 3 | 438.7376 | 90.5677
Level 4 | Candesartan | 5319.6187 | 3 | 4735.228 | 89.0144
Level 5 | Candesartan | 490.8121 | 3 | 458.6285 | 93.4428
Level 4 | Carbamazepine | 4934.7558 | 3 | 4669.2337 | 94.6193
Level 5 | Carbamazepine | 503.1905 | 3 | 469.9429 | 93.3926
Level 4 | Citalopram | 5316.9152 | 3 | 5213.1862 | 98.0491
Level 5 | Citalopram | 507.3111 | 3 | 456.5003 | 89.9843
Level 4 | Clarithromycin | 5440.9016 | 3 | 5394.9323 | 99.1551
Level 5 | Clarithromycin | 488.3816 | 3 | 522.7627 | 107.0398
Level 4 | Diclofenac | 26954.5008 | 3 | 28384.9461 | 105.3069
Level 5 | Diclofenac | 2544.3094 | 3 | 2430.8989 | 95.5426
Level 4 | Irbesartan | 5364.5547 | 3 | 4873.1565 | 90.8399
Level 5 | Irbesartan | 438.528 | 3 | 467.6877 | 106.6495
Level 4 | Metoprolol | 5246.6424 | 3 | 4727.1066 | 90.0977
Level 5 | Metoprolol | 497.8985 | 3 | 442.5561 | 88.8848
Level 4 | Propranolol | 5252.2486 | 3 | 4700.2564 | 89.4904
Level 5 | Propranolol | 519.4531 | 3 | 479.1442 | 92.2401
Level 4 | Tramadol | 5042.927 | 3 | 4667.2961 | 92.5513
Level 5 | Tramadol | 480.8009 | 3 | 460.2312 | 95.7218
Level 4 | Valsartan | 5271.0027 | 3 | 5026.0554 | 95.3529
Level 5 | Valsartan | 475.5076 | 3 | 491.0223 | 103.2628
Level 4 | Venlafaxine | 5284.0853 | 3 | 4716.1778 | 89.2525
Level 5 | Venlafaxine | 506.3362 | 3 | 457.7995 | 90.4141
SI14: working solution storage at 4 degree celsius for 14 days
sample_level | compound | conc_1 | n | conc_2 | recovery_percent
Level 4 | Amisulpride | 5016.3807 | 3 | 4366.5431 | 87.0457
Level 5 | Amisulpride | 484.4306 | 3 | 438.3335 | 90.4843
Level 4 | Candesartan | 5319.6187 | 3 | 4594.6838 | 86.3724
Level 5 | Candesartan | 490.8121 | 3 | 434.4442 | 88.5154
Level 4 | Carbamazepine | 4934.7558 | 3 | 4796.332 | 97.1949
Level 5 | Carbamazepine | 503.1905 | 3 | 449.896 | 89.4087
Level 4 | Citalopram | 5316.9152 | 3 | 5003.9885 | 94.1145
Level 5 | Citalopram | 507.3111 | 3 | 439.9605 | 86.724
Level 4 | Clarithromycin | 5440.9016 | 3 | 5585.7463 | 102.6621
Level 5 | Clarithromycin | 488.3816 | 3 | 501.3955 | 102.6647
Level 4 | Diclofenac | 26954.5008 | 3 | 30069.1066 | 111.555
Level 5 | Diclofenac | 2544.3094 | 3 | 2421.9799 | 95.192
Level 4 | Irbesartan | 5364.5547 | 3 | 4653.7042 | 86.7491
Level 5 | Irbesartan | 438.528 | 3 | 449.8696 | 102.5863
Level 4 | Metoprolol | 5246.6424 | 3 | 4581.6827 | 87.326
Level 5 | Metoprolol | 497.8985 | 3 | 447.1525 | 89.808
Level 4 | Propranolol | 5252.2486 | 3 | 4672.0737 | 88.9538
Level 5 | Propranolol | 519.4531 | 3 | 475.6796 | 91.5732
Level 4 | Tramadol | 5042.927 | 3 | 4651.0789 | 92.2297
Level 5 | Tramadol | 480.8009 | 3 | 442.7603 | 92.0881
Level 4 | Valsartan | 5271.0027 | 3 | 4795.2142 | 90.9735
Level 5 | Valsartan | 475.5076 | 3 | 503.1794 | 105.8194
Level 4 | Venlafaxine | 5284.0853 | 3 | 4665.3046 | 88.2897
Level 5 | Venlafaxine | 506.3362 | 3 | 449.4927 | 88.7736

## carryover
SI-15: carryover
compound | mean_carryover | n
Amisulpride | 0.012999 | 6
Candesartan | 0.02259 | 6
Carbamazepine | 0.018563 | 6
Citalopram | 0.013059 | 6
Clarithromycin | 0.0054266 | 6
Diclofenac | 0.056087 | 6
Irbesartan | 0.032052 | 6
Metoprolol | 0.022092 | 6
Propranolol | 0.016045 | 6
Tramadol | 0.0068929 | 6
Valsartan | 0.01307 | 6
Venlafaxine | 0.02245 | 6